## Slide 1
KCLG MEDICAL EDUCATION SERIES
Companion Diagnostics
in Oncology
Regulatory Framework · Analytical & Clinical Validation · CDx Co-Development
NGS Platforms · Liquid Biopsy · Landmark Case Studies
KCLEAGENICS MEDICAL INC. · kclgmedical.com · July 2026
Notes: Welcome to Companion Diagnostics in Oncology — a comprehensive review at PhD level. This deck covers the full CDx landscape from regulatory framework through to emerging frontiers.

## Slide 2
FOUNDATIONS
What Is a Companion Diagnostic?
Definition (FDA, 21 CFR Part 809)
"An in vitro diagnostic (IVD) device or imaging tool that provides information ESSENTIAL for the safe and effective use of a corresponding therapeutic product."
Key word: ESSENTIAL
→ Drug cannot be safely prescribed
without a positive CDx result
→ CDx result required in prescribing
information (drug label)
Companion Diagnostic
Result REQUIRED before drug use
Regulated as Class III Device — PMA pathway
Must be co-developed and co-submitted with drug
Complementary Diagnostic
Result USEFUL but not mandatory
May use 510(k) or De Novo pathway
Drug can be prescribed without positive result
Notes: A CDx is legally required before the drug can be prescribed. This is the "essential" test — contrasted with complementary diagnostics where testing is optional.

## Slide 3
MARKET OVERVIEW
CDx in Numbers: A Rapidly Growing Field
50+
FDA-Approved
CDx Devices
(as of 2025)
$10B
Global CDx
Market Value
(2024)
~$19B
Projected
Market Value
(2030)
2×
Phase III Success
Rate: Biomarker-
Selected vs. All-Comers
First CDx Approval: HercepTest (1998) · First NGS Panel CDx: FoundationOne CDx (2017) · First Liquid Biopsy CDx: cobas EGFR Plasma (2016)
Sources: FDA CDx List (updated Feb 2026); Grand View Research CDx Market Report 2024; DiMasi et al., Biomarker-selected trial success rates
Notes: The CDx market is one of the fastest-growing segments in diagnostics, driven directly by the pipeline of biomarker-defined oncology drugs. ~18% CAGR through 2030.

## Slide 4
REGULATORY FRAMEWORK
FDA Regulation: PMA Under 21 CFR Part 809
01
Class III Device
CDx classified as highest-risk device. Requires Premarket Approval (PMA) — most rigorous FDA device pathway. Regulated by CDRH.
02
Co-Development
Drug (CDER) and CDx (CDRH) developed and reviewed simultaneously. Cross-agency coordination required. Joint advisory committees in some cases.
03
PMA Submission
Dossier includes: analytical performance data, device description, software validation (for NGS), clinical data showing biomarker-by-treatment interaction.
04
Co-Approval
CDx approved contemporaneously with drug NDA/BLA. CDx result required in prescribing information before patient receives drug.
Notes: Class III designation means CDx carries the highest risk classification — paralleling the drug's potential to harm patients if given to the wrong population. PMA requires human clinical data, unlike 510(k).

## Slide 5
CDX SCIENCE
Analytical Validation: Proving the Device Works
Parameter
Definition
Oncology Implication
Diagnostic Sensitivity
% of truly positive specimens correctly identified as positive
Low sensitivity → eligible patients denied treatment. FDA scrutinises sensitivity in enrichment CDx.
Diagnostic Specificity
% of truly negative specimens correctly identified as negative
Low specificity → biomarker-negative patients receive drug without evidence of benefit — toxicity without gain.
Precision: Repeatability
Reproducibility same operator, instrument, day
Within-run variation must not cross patient-classification threshold (e.g., IHC 3+ boundary).
Precision: Reproducibility
Across operators, instruments, labs, days
Critical for IHC (subjective scoring) and NGS (bioinformatics pipeline variation). Multi-site data required.
Notes: Analytical validation answers: does this device accurately and reproducibly measure the biomarker? Without this foundation, clinical validity data is uninterpretable.

## Slide 6
CDX SCIENCE
Clinical Validation: Proving the Biomarker Predicts Benefit
The Biomarker-by-Treatment Interaction
Clinical validity requires a STATISTICALLY SIGNIFICANT INTERACTION between biomarker status and treatment effect.
This is NOT the same as showing biomarker-positive patients have better outcomes (that is a prognostic biomarker).
Predictive = differential treatment benefit by status
Prognostic = outcome regardless of treatment
The Cutoff Selection Problem
Many biomarkers are continuous variables:
· PD-L1 Tumour Proportion Score (TPS) 0–100%
· HER2 copy number ratio
· TMB (mutations per megabase)
Cutoff too HIGH:
Excludes patients who would benefit — reduces
commercial reach; denies therapy to patients.
Cutoff too LOW:
Includes non-responders — dilutes observed
efficacy signal; misleading regulatory data.
Notes: Clinical validity is the hardest part. It requires trial data showing the biomarker-by-treatment interaction — not just that positive patients do better, but that they benefit MORE from the drug than negative patients.

## Slide 7
CDX DEVELOPMENT
CDx Co-Development Timeline
Pre-IND
Drug
Identify molecular target; biomarker hypothesis from preclinical data
CDx
Candidate biomarker assay selection; CDx partner identification; FDA pre-IND biomarker discussion
Phase I
Drug
Safety/PK; dose escalation; biomarker expansion cohort if early signals emerge
CDx
Evaluate candidate assay on Phase I specimens; begin analytical development; feasibility testing
Phase II
Drug
Initial efficacy; define biomarker-positive population; proof-of-concept
CDx
Lock prototype CDx assay; complete analytical performance package; begin CDRH pre-submission meetings
Phase III
Drug
Randomised confirmatory trial; CDx used as inclusion criterion
CDx
Commercial-intent assay in pivotal trial; accumulate PMA data; submit PMA parallel with NDA/BLA
Notes: Late-stage CDx development is the most common cause of drug-CDx approval misalignment. Early alignment — from Phase I onward — is essential for contemporaneous approval.

## Slide 8
CASE STUDY 1 / HER2 & TRASTUZUMAB
The Original CDx: HercepTest & Herceptin (1998–2024)
Timeline of the HER2/CDx Story
· 1998 — HercepTest (IHC, Dako) approved as first-ever FDA CDx alongside trastuzumab (Herceptin) for HER2+ metastatic breast cancer
· 2001 — Slamon et al. (NEJM): trastuzumab + chemo improved OS 20.3 → 25.1 months (p=0.046) in IHC 3+ or FISH+ patients
· 2007 — ASCO/CAP guidelines revealed up to 20% of "HER2-positive" patients may have been misclassified — wrong treatment, wrong denial
· 2012 — Multiple HER2 assay platforms (IHC 3+, FISH ratio ≥2.0, ISH) in simultaneous use across countries with different cutoffs
· 2017 — FoundationOne CDx approved as pan-tumour NGS CDx for HER2 (ERBB2) amplification
· 2023 — ASCO/CAP update added HER2-low category (IHC 1+, or IHC 2+/ISH-) for trastuzumab deruxtecan (T-DXd) — a new patient population invisible to the original binary framework
Biomarker
ERBB2 amplification
or protein overexpression
CDx Devices
HercepTest IHC · PathVysion FISH · FoundationOne CDx
Drug (NDA/BLA)
Trastuzumab BLA 103792
Plus 10+ HER2-targeted agents
Key Lesson
CDx complexity can persist for decades. New CDx categories can open new patient populations 25 years post-approval.
Notes: The HER2/trastuzumab story is the most instructive in CDx history. It illustrates how CDx complexity can accumulate over time, how misclassification at scale affects patients, and how a new CDx category (HER2-low) can open a previously invisible patient population decades after initial approval.

## Slide 9
CASE STUDY 2 / EGFR & OSIMERTINIB
Sequential CDx Development: 1st to 3rd Generation EGFR TKI
1st Gen
Drug
Erlotinib
Gefitinib
CDx
cobas EGFR v1
(tissue: Ex19del, L858R)
Outcome
PFS ~9-11 mo
All patients develop resistance
2nd Gen
Drug
Afatinib
CDx
cobas EGFR v2
(tissue: Ex19del, L858R)
Outcome
PFS ~11-13 mo
~60% acquire T790M
T790M
CDx
Drug
Plasma CDx
First in Class
CDx
cobas EGFR v2
PLASMA — T790M
(2016, first liquid biopsy CDx)
Outcome
Plasma sensitivity:
~70-75% vs tissue gold standard
3rd Gen
Drug
Osimertinib
(Tagrisso)
CDx
cobas EGFR v2 (tissue + plasma)
FoundationOne CDx
Guardant360 CDx (plasma)
Outcome
FLAURA: PFS 18.9 vs 10.2 mo
HR 0.46 (p<0.001)
Gold standard 1L
Notes: The EGFR/osimertinib case is the best example of sequential CDx development. Each therapeutic generation required a new CDx — and the T790M CDx in plasma was a watershed moment, establishing liquid biopsy as a viable CDx platform.

## Slide 10
CASE STUDY 3 / BRCA & PARP INHIBITORS
Germline vs. Somatic CDx: BRACAnalysis CDx and FoundationOne CDx
Olaparib (Lynparza) CDx Approvals
2014
Ovarian Cancer
BRACAnalysis CDx (germline) | BRCA1/2 mutations
2018
Breast Cancer
BRACAnalysis CDx (germline) | BRCA1/2 mutations
2019
Pancreatic Cancer
BRACAnalysis CDx (germline) | BRCA1/2 mutations
2020
mCRPC (Prostate)
BRACAnalysis CDx + FoundationOne CDx | BRCA1/2 (germline)
HRR panel (somatic — 14 genes)
Germline (BRACAnalysis CDx)
Whole blood DNA — detects inherited BRCA1/2 variants. First CDx based on germline testing. Submitted by Myriad Genetic Laboratories. Approved 2014 (P140020).
Somatic (FoundationOne CDx)
Tumour tissue DNA — detects somatic BRCA1/2 + all HRR gene alterations (14 genes in mCRPC). Enables treatment for the ~5-7% with somatic-only BRCA mutations. Same drug (olaparib), different CDx, different patient population.
Notes: The BRCA/PARP inhibitor landscape shows how a single biomarker can drive approvals across multiple tumour types, and how germline and somatic testing create parallel CDx pathways.The BRCA CDx landscape is the clearest example of how a biomarker evolves across indications. Note the addition of the full HRR panel (14 genes) for mCRPC — not just BRCA1/2 — which required a different CDx device (FoundationOne CDx) for the same drug.

## Slide 11
CASE STUDY 4 / PD-L1 & PEMBROLIZUMAB
Continuous Biomarker Complexity: PD-L1 IHC and TMB
Indication
Required CDx
Threshold / Note
NSCLC 1L Monotherapy
PD-L1 IHC 22C3 pharmDx
TPS ≥50% (~30% of NSCLC qualify)
NSCLC 1L + Chemo Combo
PD-L1 IHC 22C3 pharmDx
TPS ≥1%
Cervical Cancer 2L Mono
PD-L1 IHC 22C3 pharmDx
TPS ≥1%
MSI-H / dMMR Tumours
(any histology)
MMR IHC (local) or validated PCR
No PD-L1 test required — histology-agnostic
TMB-High Solid Tumours
(any histology)
FoundationOne CDx
TMB ≥10 mut/Mb — first pan-tumour genomic CDx (2020)
Key insight: For the same drug, 5 different testing approaches — different assays, different thresholds, different biological rationales.
Notes: Pembrolizumab has the most complex CDx profile in oncology — because the required test and threshold differ by tumour type, line of therapy, and whether it is mono or combo therapy.The PD-L1 pembrolizumab case is the most operationally complex CDx scenario in oncology today. Site staff and CRAs must know which CDx and which threshold applies to the specific pembrolizumab indication being studied — these are not interchangeable.

## Slide 12
NGS PLATFORM CDX
FoundationOne CDx: The Pan-Tumour NGS Paradigm
324
cancer-related genes
analysed in one test
+ TMB status
+ MSI status
+ Genomic signatures
Approved: Nov 2017 (P170019)
Foundation Medicine / Roche
EGFR Ex19del / L858R
NSCLC
Osimertinib, Erlotinib, Gefitinib, Afatinib
ALK rearrangements
NSCLC
Alectinib, Crizotinib, Ceritinib
BRAF V600E
NSCLC, Melanoma
Dabrafenib + Trametinib
ERBB2 (HER2) amplification
Breast, Gastric
Trastuzumab, Pertuzumab, T-DM1
BRCA1/2 alterations
Multi-tumour
Olaparib (ovarian, breast, prostate, pancreatic)
PIK3CA mutations
HR+/HER2- Breast
Alpelisib (Piqray)
MET exon 14 skipping
NSCLC
Capmatinib (Tabrecta)
TMB ≥10 mut/Mb
All solid tumours
Pembrolizumab (Keytruda)
Notes: FoundationOne CDx (P170019, approved Nov 2017) transformed the CDx landscape by consolidating multiple single-analyte assays into a single pan-tumour NGS panel that can serve as CDx for multiple drugs simultaneously.FoundationOne CDx approval fundamentally changed CDx economics and logistics. A single tissue sample can now determine eligibility for multiple drugs across multiple biomarkers. The approval process for each new biomarker/drug pair is a PMA supplement (S-number), building on the original P170019 approval.

## Slide 13
LIQUID BIOPSY
ctDNA Liquid Biopsy as Companion Diagnostic
How ctDNA Liquid Biopsy CDx Works
Tumour cells shed DNA fragments into blood circulation (ctDNA). A blood draw detects tumour mutations without tissue biopsy.
FDA-APPROVED LIQUID BIOPSY CDx:
Guardant360 CDx (Guardant Health)
EGFR Ex19del / L858R in plasma
For 1st-line osimertinib in NSCLC
cobas EGFR Mutation Test v2 (Roche)
EGFR T790M in plasma
For 2nd-line osimertinib post-TKI
Advantages
· Minimally invasive (blood draw)
· Captures tumour heterogeneity from multiple sites
· Repeatable — tracks mutation evolution during treatment
· Valuable when tissue is insufficient, inaccessible, or high-risk
Key Limitations
· Reduced sensitivity in early-stage disease (low shedding)
· False-negative rate up to 30% in stage I–II NSCLC
· Fusions/amplifications less reliably detected in plasma
· NEGATIVE result does NOT exclude mutation
→ Tissue biopsy must follow (plasma-first, tissue-reflex)
Notes: Liquid biopsy CDx eliminates the need for tissue biopsy in appropriate settings. Key limitation: a negative plasma result does NOT exclude the mutation. Always specify a tissue-reflex algorithm in clinical trials.Critical operational point for trial teams: a negative Guardant360 CDx result in plasma does NOT make a patient ineligible. Protocol must specify reflex to tissue biopsy CDx. A positive plasma result IS sufficient for eligibility in the approved osimertinib indications.

## Slide 14
TRIAL DESIGN IMPLICATIONS
Biomarker Trial Design: Enrichment, Stratification & Interaction
Enrichment Design
When to Use
Strong prior evidence of differential benefit in biomarker+ patients; rare biomarker requiring central testing
Regulatory
Requires prospective validated CDx at screening; FDA expects parallel CDx development pathway
Advantages
Smaller sample size; more efficient efficacy detection; less exposure of likely-non-responders to drug toxicity
Limitations
Cannot formally test biomarker-by-treatment interaction; limits label to tested population only
All-Comers with Stratification
When to Use
Hypothesised differential benefit but insufficient evidence for enrichment; biomarker prevalence high enough to enrol both groups
Regulatory
Pre-specify biomarker subgroup in SAP; powered on overall population; subgroup hypothesis-generating unless powered
Advantages
Can generate interaction test data; broader eligibility; pragmatic intent-to-treat effect estimable
Limitations
Requires larger sample size; biomarker-negative patients receive drug without expected benefit
Biomarker-Stratified Design
When to Use
Genuine equipoise about benefit in biomarker-negative patients; want formal test of interaction
Regulatory
Complex multiplicity adjustment required; hierarchical testing procedures specified in protocol
Advantages
Full interaction test powered; potential for two separate label claims (BM+ and BM-)
Limitations
Largest sample size requirement; most complex statistical plan; less common in modern oncology
Notes: The CDx strategy determines trial design. Enrichment designs are operationally simpler but cannot formally demonstrate the biomarker-by-treatment interaction required for CDx clinical validation. All-comers designs can test the interaction but require larger sample sizes.Slide for PIs and regulatory colleagues. The choice of enrichment vs. all-comers design determines not just the statistical plan but the type of evidence generated for CDx validation. Enrichment designs cannot provide the interaction test data needed for standalone CDx clinical validation.

## Slide 15
GLOBAL REGULATION
CDx Regulation Beyond FDA: EMA, PMDA, NMPA
United States (FDA)
21 CFR Part 809 · Class III · PMA pathway · CDRH review · Contemporaneous approval with drug (CDER) · ~600 day review target
Most established CDx framework globally
European Union (EMA/IVDR)
EU IVDR 2017/746 · Full effect from 2022 · Notified Body review (not direct EMA) · Higher evidence standard than prior IVDD · Companion diagnostics designated as Class C or D devices
Significant regulatory step-change from old IVDD
Japan (PMDA)
Simultaneous regulatory review (SAKIGAKE) introduced 2014 · CDx reviewed by PMDA alongside drug · Bridging study often required for assay equivalence · Japanese-specific population data may be required
Alignment with FDA/EMA increasing but Japan-specific data common
China (NMPA)
CDx regulated under medical device framework · Domestic clinical data often required · Increasing alignment with ICH guidelines since 2017 ICH membership · Complex regulatory environment for global submissions
Rapidly evolving; local clinical data requirement remains key barrier
Notes: Global CDx development must account for parallel regulatory pathways that may differ substantially in evidence requirements and approval timelines. A US-approved CDx does not automatically satisfy European, Japanese, or Chinese requirements.For global oncology trials, the CDx regulatory strategy must account for multiple simultaneous pathways. The EU IVDR transition (2022) was particularly impactful — previously CE-marked IVDs required full re-evaluation under the new framework with substantially higher evidence requirements.

## Slide 16
CLINICAL TRIAL OPERATIONS
CDx Readiness Checklist for Trial Teams
HIGH
CDx vs. trial assay alignment
Assay used for enrolment = commercial CDx or validated as equivalent by bridging study
HIGH
Local vs. central testing defined
Protocol specifies which result governs eligibility; both cannot override the other without protocol definition
HIGH
Pre-analytical requirements trained
Archival tissue age, fixation duration, tumour content %, fresh biopsy criteria — all site staff trained
MED
Turnaround time compliance
CDx TAT consistent with screening window; contingency procedure for exceedances documented
HIGH
Negative result algorithm
Plasma ctDNA negative → tissue reflex procedure specified; eligibility determination documented
MED
Screen failure documentation
CDx result recorded for all screen failures — essential for biomarker prevalence estimation and protocol review
MED
Tissue banking and consent coverage
Residual CDx specimens banked; consent covers future biomarker research; chain of custody maintained
HIGH
Regulatory alignment confirmed
Pre-IND/Type B meeting confirms FDA accepts the proposed CDx for this trial; documented in IND
Notes: This checklist is directly applicable to site initiation, monitoring visits, and protocol deviation prevention. CDx-related screen failures and eligibility errors are among the most common FDA inspection findings in biomarker-enriched trials.This checklist is adapted from FDA inspection findings in biomarker-enriched oncology trials. The most common CDx-related deficiencies in FDA site inspections: using an unvalidated local assay for eligibility determination, missing screen failure CDx documentation, and inadequate tissue pre-analytical training.

## Slide 17
EMERGING FRONTIERS
Beyond Today's CDx: Where the Field Is Going
01
Multi-Omic CDx
Integrating genomic, transcriptomic, proteomic, and epigenomic biomarkers into a single predictive algorithm. FDA's AI/ML Device Action Plan provides an emerging framework for algorithm-based CDx approval.
02
MRD as CDx Endpoint
Minimal residual disease (MRD) detection as a CDx endpoint for adjuvant therapy decisions. FDA draft guidance (2020) on MRD in haematologic malignancies; solid tumour applications in development.
03
AI-Powered Digital Pathology
AI image analysis of standard H&E slides to extract biomarker information without additional staining or molecular testing. Multiple pre-submissions with FDA ongoing; spatial transcriptomics integration next step.
04
ctDNA Expansion
ctDNA CDx approval expanding beyond EGFR/NSCLC to BRAF, KRAS, and pan-cancer biomarkers. Tumour-informed ctDNA MRD assays (personalised) creating new approval pathway discussions with FDA.
Notes: These frontiers represent the next wave of CDx development. AI-based digital pathology CDx is currently in FDA pre-submission discussions. MRD as a regulatory endpoint has draft FDA guidance in haematologic malignancies.

## Slide 18
KEY TAKEAWAYS
Companion Diagnostics: What Every Oncology
Researcher Needs to Know
1
CDx ≠ optional testing. The word "essential" in the FDA definition has legal force — a drug cannot be prescribed without a positive CDx result.
2
Analytical validity and clinical validity are both required. A device can accurately detect a mutation (analytical) but fail if that mutation doesn't predict differential benefit (clinical).
3
Simultaneous co-development is not optional. Late-stage CDx development is the most common cause of drug-CDx approval misalignment — start the CDx programme in Phase I.
4
NGS panels (FoundationOne CDx) can serve as CDx for multiple drugs simultaneously — but each drug/biomarker combination requires its own PMA supplement approval.
5
A negative liquid biopsy result does NOT exclude the mutation. Every protocol using plasma CDx must specify a tissue-reflex algorithm.
6
PD-L1 and TMB are not interchangeable biomarkers. The required assay, threshold, and clinical rationale differ by indication — even within pembrolizumab alone.
Notes: Summary slide for closing. Emphasise the co-development point — it is the single biggest operational lever. Companies that start CDx development late consistently face delays at submission. Also emphasise the liquid biopsy caveat — it is a frequent protocol deviation trigger in the field.

## Slide 19
REFERENCES
References & Regulatory Sources
1. FDA. In Vitro Companion Diagnostic Devices — Guidance for Industry and Food and Drug Administration Staff. August 2014 (updated 2016). FDA.gov
2. FDA. List of Cleared or Approved Companion Diagnostic Devices (IVD and Imaging Tools). fda.gov/medical-devices/in-vitro-diagnostics. Updated February 2026.
3. 21 CFR Part 809 — In Vitro Diagnostics General Regulations. FDA Code of Federal Regulations.
4. EU IVDR 2017/746 — Regulation on in vitro diagnostic medical devices. European Parliament and Council, 2017.
5. Slamon DJ et al. Use of Chemotherapy plus a Monoclonal Antibody against HER2 for Metastatic Breast Cancer. NEJM. 2001;344(11):783–792.
6. Ramalingam SS et al. (FLAURA). Overall Survival with Osimertinib in Untreated, EGFR-Mutated Advanced NSCLC. NEJM. 2020;382:41–50.
7. Mok TS et al. (PACIFIC extended data). Osimertinib or Platinum-Pemetrexed in p.T790M-Positive Lung Cancer. NEJM. 2017;376:629–640.
8. DiMasi JA et al. Success Rates for New Drugs Entering Clinical Testing in the United States. Clin Pharmacol Ther. 2010;87(3):272–277.
9. Foundation Medicine FoundationOne CDx PMA P170019 (and supplements). FDA CDRH Database.
10. Guardant Health. Guardant360 CDx. PMA P200010. FDA Approval 2022.
11. Myriad Genetics BRACAnalysis CDx. PMA P140020 (and supplements). FDA CDRH Database.
12. Wolchok JD et al. iRECIST: Guidelines for response criteria for use in trials testing immunotherapeutics. Lancet Oncol. 2017;18(3):e143–e152.
13. Merck & Co. Keytruda (pembrolizumab) Prescribing Information. 2024. FDA-approved labelling.
14. FDA. Project Optimus: Dose Optimization in Oncology Drug Development. FDA.gov/oncology. 2022.
Notes: All regulatory sources are publicly available on FDA.gov and ICH.org. The CDx list is updated regularly by FDA CDRH.

## Slide 20
KCLEAGENICS MEDICAL INC.
GCP-Certified CRO · Oak Brook, Illinois
Oncology · Haematology · Metabolic Medicine
www.kclgmedical.com
keog@kclgmedical.com
312-225-3916
GCP Certified
FDA 21 CFR Compliant
ICH GCP E6(R3) Aligned
ClinicalTrials.gov Registered
Learn More
Education Hub
kclgmedical.com/education.html
Related Articles:
· Companion Diagnostics in Oncology (Full Article)
· Protocol Preparation in Oncology Trials
· Risk-Based Monitoring in Oncology
· ICH GCP E6(R3): Guide for Oncology Sites
· Biospecimen Handling in Oncology Trials